## Failed PTs and Well Info
Well | Failed PT SN | Failed PT Install Date | Failed PT Removal Date | Failed PT Cal Due | PT Failure Date | DTW (at time of failure/ removal) | Usual PSI | Cable SN | Cable Install Date | Cable Measured Length | PT Install Depth | Surface Check | Passed Surface Check? | Failure Type | Telemetry Installed? (at time of failure) | Telemetry Install Date | Notes
ERDA-9 | 21830041 | 2021-07-21 | 2021-03-19 | 2021-06-11 | 2021-01-18 | 409.48 | 156.272 | 215667 | 2008-11-04 | 759.3 | 717.20 | 24.316 | No | Catastrophic | No | N/A | PSI measurements were abnormal. Same issue occurred on 7/21/2020 with PT S/N: 21237051
SNL-2 | 08203061 | 2021-03-02 | 2021-04-15 | 2021-12-16 | Unknown | 263.78 | 103.405 | 200960 | 2008-11-12 | 503.31 | 470.70 | N/A | N/A | Connection Failure | Yes | 2020-12-14 | PT Failure. Possible bad (old) cable
H-11b2 | 21237081 | 2020-07-07 | 2021-05-05 | 2021-06-11 | 2021-04-23 | 265.58 | 180.434 | 221571 | 2016-07-27 | 717.4 | 634.91 | 10.456 | No | Catastrophic | No | N/A | PSI measurements were abnormal. Same issue occurred on 7/7/2020 with PT S/N: 21534016
H-6Br | 21918013 | 2021-03-03 | 2021-05-05 | 2022-02-25 | N/A | 299.24 | 155.930 | 476362 | 2018-02-15 | 659.2 | 616.60 | 13.156 | Yes | Concerning psi readings | No | 2021-05-28 | No failure, psi was concerning so it was pulled.
SNL-2 | 21741069 | 2021-04-15 | 2021-05-19 | 2022-03-11 | N/A | 263.65 | 103.225 | 200960 | 2008-11-12 | 503.31 | 470.70 | 13.151 | Yes | Low Battery | Yes | 2020-12-14 | PT was about to fail due to 10% battery.
WQSP-2 | 21741067 | 2020-08-27 | 2021-06-10 | 2021-08-24 | Unknown | N/A | 193.640 | 634483/632396 | 2019-02-14 | 819/5 | 814.00 | N/A | N/A | Connection Failure | Yes | 2019-10-01 | PT Failure dicovered on 5/18/21.
WQSP-3 | 03200500 | 2020-08-28 | 2021-06-10 | 2021-08-24 | Unknown | N/A | 198.890 | 634481/632372 | 2019-02-14 | 854/5 | 849.00 | N/A | N/A | Connection Failure | Yes | 2019-10-01 | PT Failure dicovered on 5/18/21.
SNL-14 | 21534015 | 2021-05-05 | 2021-06-16 | 2022-03-17 | Unknown | 390.86 | 139.191 | 212112 | 2008-09-24 | 712.52 | 668.95 | N/A | N/A | Connection Failure | Yes | 2020-04-28 | PT Failure. Possible bad (old) cable
SNL-2 | 21830024 | 2021-05-19 | 2021-06-16 | 2022-03-16 | 2021-05-27 | 263.74 | 103.596 | 200960 | 2008-11-12 | 503.31 | 470.70 | 332.322 | No | Low Battery&Catastrophic | Yes | 2020-12-14 | Bad (old) cable. Cable has been replaced on 6/21/21 and downloads have been successful.
C-2737c | 21237072 | 2021-03-02 | 2021-07-02 | 2021-07-02 | 2021-06-12 | 398.99 | 142.613 | 606871 | 2019-06-21 | 711.35 | 688.85 | 14.145 | No | Concerning psi readings | Yes | 2018-09-01 | PSI measurements were abnormal.
SNL-5 | 21237059 | 2020-07-16 | 2021-07-06 | 2021-07-08 | Unknown | 319.88 | 156.404 | 597741 | 2019-02-04 | 700.00 | 649.00 | N/A | N/A | Connection Failure | Yes | 2020-10-20 | PT Failure. One of two AA batteries inverted polarity.
SNL-14 | 21927002 | 2021-06-16 | 2021-07-13 | 2022-03-16 | Unknown | 390.48 | 139.77 | 212112 | 2008-09-24 | 712.52 | 668.95 | N/A | N/A | Connection Failure | Yes | 2020-04-28 | PT Failure. Possible bad (old) cable
AEC-7R | 21918012 | 2021-01-13 | 2021-07-14 | 2021-12-16 | Unknown | 616.33 | 128.436 | 268361 | 2017-05-11 | 918.89 | 866.00 | N/A | N/A | Connection Failure | Yes | 2020-08-13 | Possible bad cable or telemetry fry? Diaphragm oozing green substance. Batteries intact.
SNL-3 | 03200494 | 2020-11-04 | 2021-07-16 | 2021-11-03 | Unknown | 429.25 | 163.360 | 212604 | 2008-07-23 | 811.83 | 766.50 | N/A | N/A | Connection Failure | Yes | 2020-08-21 | PT Failure. Bad (old) cable
SNL-17A | 21237007 | 2020-09-14 | 2021-07-27 | 2021-09-10 | Unknown |  | 61.999 | 200924 | 2007-01-25 | 407.63 | 349.60 | N/A | N/A | Connection Failure | Yes | 2021-07-06 | PT Failure. Possible bad (old) cable
SNL-9 | 08200111 | 2021-02-03 | 2021-08-17 | 2022-02-02 | Unknown | 319.63 | 123.583 | 606863 | 2008-10-21 | 604.78 | 567.20 | N/A | N/A | Connection Failure | Yes | 2021-03-04 | PT Failure. Possible bad (old) cable
SNL-8 | 21933001 | 2021-03-02 | 2021-08-18 | 2022-02-25 | Unknown | 542.61 | 214.722 | 268351 | 2016-12-15 | 1007.82 | 969.70 | N/A | N/A | Connection Failure | Yes | 2021-01-28 | PT Failure
SNL-5 | 03200499 | 2021-07-06 | 2021-08-18 | 2022-06-28 | Unknown | 319.26 | 156.404 | 597741 | 2019-02-04 | 700.00 | 649.00 | N/A | N/A | Connection Failure | Yes | 2020-10-20 | PT Failure
AEC-7R | 03200493 | 2021-07-14 | 2021-08-18 | 2022-06-29 | Unknown | 616.16 | 128.536 | 268361 | 2017-05-11 | 918.89 | 866.00 | N/A | N/A | Connection Failure | Yes | 2020-08-13 | PT Failure
SNL-15 | 08200112 | 2021-05-18 | 2021-08-19 | 2022-05-10 | Unknown | 473.14 | 252.216 | 268349 | 2017-05-25 | 1000 | 922.18 | N/A | N/A | Connection Failure | Yes | 2021-01-28 | PT Failure
WQSP-3 | 06202172 | 2021-06-10 | 2021-08-25 | 2022-06-08 | Unknown | N/A | 198.890 | 634481/632372 | 2019-02-14 | 854/5 | 849.00 | N/A | N/A | Connection Failure | Yes | 2019-10-01 | PT Failure.
WQSP-2 | 06202176 | 2021-06-10 | 2021-08-25 | 2022-06-08 | Unknown | N/A | 193.640 | 634483/632396 | 2019-02-14 | 819/5 | 814.00 | N/A | N/A | Connection Failure | Yes | 2019-10-01 | PT Failure.
WQSP-1 | 06202186 | 2021-06-10 | 2021-08-25 | 2021-06-08 | Unknown | N/A | 164.831 | 634478/632283 | 2019-02-14 | 712/5 | 707.00 | N/A | N/A | Connection Failure | Yes | 2019-10-01 | WQSP-1,2,3 all had no communication from PT to laptop during July MMR. All three PTs lasted < a month.
H-6Br | 21927001 | 2021-05-05 | 2021-09-21 | 2022-03-17 | Unknown | 299.24 | 155.780 | 476362 | 2018-02-15 | 659.2 | 616.60 | N/A | N/A | Connection Failure | Yes | 2021-05-28 | PT Failure
H-15R | 06202180 | 2021-03-02 | 2021-09-23 | 2022-01-25 | Unknown | 521.32 | 184.099 | 268358 | 2017-08-03 | 950 | 872.50 | N/A | N/A | Connection Failure | Yes | 2021-05-25 | PT Failure
SNL-14 | 06202185 | 2021-07-13 | 2021-10-19 | 2022-06-28 | Unknown | 392.29 | 138.397 | 212112 | 2008-09-24 | 712.52 | 668.95 | N/A | N/A | Connection Failure | Yes | 2021-04-28 | PT Failure. Possible bad (old) cable. Discovered no comm. 10/13/21
H-15R | 21237035 | 2021-09-23 | 2021-10-19 | 2022-09-20 | Unknown | 520.99 | 184.278 | 268358 | 2017-08-03 | 950 | 872.50 | N/A | N/A | Connection Failure | Yes | 2021-05-25 | PT Failure. Discovered no comm. 10/13/21
H-11b2 | 21830025 | 2021-06-07 | 2021-11-11 | 2022-04-27 | 2021-11-06 | 266.04 | 180.395 | 221571 | 2016-07-27 | 717.40 | 634.91 | 15.627 | No | Concerning psi readings | No | N/A | PSI measurements were abnormal.
WIPP-11R | 08200114 | 2021-09-22 | 2021-11-11 | 2022-08-24 | Unknown | 372.31 | 227.384 | 268360 | 2020-06-02 | 913.81 | 853.14 | N/A | N/A | Connection Failure | Yes | 2021-10-06 | PT Failure. Connection between PT cable and Telemetry was under water in between casing and environmental casing. Needs new PT and cable.

## 2021 New Cable Installs
Well | PT Install Depth | Old Cable SN | Old Cable Install Date | Old Cable Measured Length | Date of old Cable Pulled/New Cable Installed | New Cable Length | New Cable Measured Length | New Cable SN | Failed PT SN
SNL-2 | 470.70 | 200960 | 2008-11-12 | 503.31 |  |  |  |  | 0820361 | Cable Already replaced
SNL-3 | 766.50 | 212604 | 2008-07-23 | 811.83 |  | 775 |  |  | 03200494
SNL-14 | 668.95 | 212112 | 2008-09-24 | 712.52 |  | 712 |  |  | 21534015
SNL-17a | 349.60 | 200924 | 2007-01-25 | 407.63 |  | 407 |  |  | 21237007
SNL-9 | 567.20 | 200880 | 2008-10-21 | 604.78 |  | 600 or 575

## New Cable Inventory
Length (ft) | Quantity
375 | 1
425 | 1
450 | 1
525 | 1
550 | 1
550 | 2
575 | 1
600 | 2
625 | 3
675 | 2
700 | 2
725 | 2
750 | 1
775 | 3
900 | 0
950 | 1
1400 | 1